## Slide 1
PPV40V Bench Experiments
PWM–voltage calibration and fake-lung inspiration modelling
2026-10-06
Two experiments
1. PCB/Arduino PWM → voltage
2. PPV voltage → inspiratory flow
Preliminary parameter identification and test-quality review

## Slide 2
Introduction
Electrical command conversion and pneumatic system response
2
Experiment 1: PWM–voltage sweep
PWM
0%-100%
PCB+STM32
bias from nominal value
Multimeter
0–5 V
• Linearity, steady-state error
Experiment 2: fake-lung inspiration
PPV voltage
0–4.2 V
Open ATM
Inspiration step
SFM3300
flow rate
• Time constant, control range, closure voltage and γ
• PPV was preset; its actuation delay cannot be isolated
The electrical calibration supports interpreting the flow plateau as a pneumatic-system effect.

## Slide 3
Experiment 1: output voltage closely follows the PWM command
3
Method
Sweep 0–100%, then 100–0%
Measured results
• Maximum error: 7.4 mV • Max. up/down difference: 1.0 mV
Conclusion
Passing through the PCB and bridge board does not affect the ideal voltage output (i.e., the boards do not interfere with the pinch valve control).

## Slide 4
Experiment 2: inspiration at a fixed PPV voltage
Manually compress the fake lung, then open ATM to initiate inspiration
4
1
Compress lung
Manual initial state
2
Close ATM/VAC
Isolate the circuit
3
Set PPV
0–4.2 V
4
Open ATM
Inspiration
5
Record flow
SFM3300, 50 Hz
- Inspiratory flow curve ------ first order system
- inspired volume------tidal volumn
- Fully open time constant τin (τeff in fully open)
- Effective time constant τeff at different voltages
- PPV control range and closure voltage
- Nonlinearity parameter γ ----modeling parameter value about the PPV

## Slide 5
flow traces
6
• 0–3.2 V: The curves decay at similar rates, indicating little valve restriction.
• 3.5–3.9 V: The flow peaks decrease, and inspiration takes longer.
• 4.1 V: The flow is much lower, and the response is significantly slower.
• 4.2 V: The flow is close to the baseline, indicating that the valve is fully closed.
• Peak flow depends on the initial compression of the fake lung

## Slide 6
First-order time-constant identification
7
Initial results
The following initial modelling parameters were obtained:
• Fully open inspiratory time constant: τin ≈ 1.9 s
• Effective control onset voltage: Ud ≈ 3.1 V
• Full closure voltage: Uc ≈ 4.2 V
• Nonlinearity parameter: γ ≈ 0.65
• Effective time constant τeff
Use τin ≈ 1.9 s as the fully open baseline; the time constant rises rapidly above 3.9 V.

## Slide 7
Limitations: only inspiration was identified
10
Identified
•basic modeling parameter:
Fully open τin and restricted-opening τeff
• PPV control range, closure voltage and γ
Still unknown
• Expiratory / vacuum-collapse time constant τout
• Isolated PPV command-to-flow delay
Control implications
• Vacuum is adjustable; τout is less urgent for the current inspiration model
• τout is still needed to assess breathing rate, I:E ratio and cycle consistency
• Identify τout at several vacuum settings
Adjustable vacuum can tune expiration speed; a complete model still needs both inspiratory and expiratory parameters.

## Slide 8
Preliminary results: usable initial model parameters
Use these estimates for initial modelling and sensitivity analysis, rather than final calibration
11
τin ≈ 1.9 s
Fully open; 1.8–2.1 s
Ud ≈ 3.1 V
Control onset
Uc = 4.2 V
Confirmed closure
γ ≈ 0.65
Sensitivity: 0.60–0.70
R² = 0.986
τ–voltage fit
Suggested initial inspiration model
τeff = τin / aeff(U)^γ
aeff=1 for U≤3.1 V; decreases linearly to 0 at 4.2 V; closed for U≥4.2 V.
Applications
• Initial feedforward model parameters
• Controller parameter sensitivity tests
• Baseline for subsequent breathing-cycle tests
Apply γ≈0.65 together with the piecewise effective opening aeff(U).

## Slide 9
Future work
12
1.
Vacuum / expiration tests
Record flow at several vacuum settings; estimate τout and collapse volume.
2.
PPV step response and delay
With ATM open and pressure drop stable, step PPV and record command and flow together.
3.
Complete breathing cycles
Run consecutive inspiration / expiration cycles; check phase switching and repeatability.

## Slide 10
Appendix: time constants by voltage
Ranges show the minimum and maximum individual fitted values at each voltage
13
Voltage / V | n | Median τ / s | τ range / s | Peak flow / SLM | Median volume / mL
0 | 1 | 1.91 | 1.91–1.91 | 51.7 | 1602
1 | 2 | 1.81 | 1.76–1.85 | 47.9 | 1569
2 | 2 | 1.81 | 1.72–1.90 | 61.7 | 1674
3 | 3 | 2.02 | 1.93–2.28 | 48.5 | 1569
3.1 | 1 | 1.76 | 1.76–1.76 | 53.0 | 1571
3.2 | 2 | 2.05 | 1.92–2.18 | 41.2 | 1528
3.3 | 2 | 2.56 | 2.19–2.92 | 39.7 | 1539
3.4 | 2 | 2.23 | 2.16–2.31 | 37.7 | 1513
3.5 | 5 | 2.39 | 2.04–2.99 | 35.5 | 1375
3.6 | 4 | 2.72 | 2.62–3.61 | 23.0 | 1396
3.7 | 2 | 3.29 | 3.25–3.33 | 16.3 | 1291
3.75 | 3 | 3.76 | 2.97–3.77 | 17.9 | 1342
3.8 | 3 | 3.44 | 3.23–3.51 | 16.5 | 1323
3.85 | 2 | 3.77 | 3.62–3.92 | 18.4 | 1078
3.9 | 2 | 4.67 | 4.62–4.72 | 12.4 | 1245
4 | 1 | 5.38 | 5.38–5.38 | 13.1 | 1193
4.1 | 1 | 9.41 | 9.41–9.41 | 5.6 | 837
4.2 | 1 | Closed | — | ≈0 | —